The following table represents examples of equipment that is battery or electrically operated and the types of servicing and calibration, and/or monitoring and checking typically encountered.

| Item | Servicing/Calibration | Monitoring/Checking | Parts/Supplies |
| --- | --- | --- | --- |
| Audiometer | In accordance with the manufacturer’s recommendations. | In accordance with the manufacturer’s recommendations. |  |
| Auriscope | Monthly clean of viewing glass or as required. | Replace or recharge batteries when light dims. | Spare batteries (standard or rechargeable), globe |
| Computer system including backups | According to the maintenance and servicing program. | Server temperature. | Backup media, spare computer |
| Dermoscope | Monthly clean of viewing glass or more frequently as required. | Replace or recharge batteries when light dims. | Spare batteries (standard or rechargeable), globe |
| Defibrillator |  | Monthly check of function or in accordance with the manufacturer’s recommendations. | Battery |
| ECG | Annual servicing and calibration by qualified technician. | Each ECG push check button (1mv=large square). | ECG paper |
| Electrical cords and all computers, printers and appliances | Annual testing and tagging by a qualified person. |  |  |
| Fire Extinguishers | Biannual check by qualified technician ie: MFB or private service. |  |  |
| Generator | Annual full test of function. | Monthly check fuel level and starts. | Fuel |
| Glucometer | Annual comparison of an actual blood sugar level with a glucometer reading of same blood sample. | Monthly. | Test strips, batteries |

| Item | Servicing/Calibration | Monitoring/Checking | Parts/Supplies |
| --- | --- | --- | --- |
| Non-medical items e.g. door mats, floor surfaces | Yearly inspection and replacement as required. | Warning signs displayed when wet or fall risk is evident. |  |
| Ophthalmoscope | Monthly clean of viewing glass or as required. | Replace or recharge batteries when light dims. | Spare batteries (standard or rechargeable), globe |
| Oxygen | Annual servicing and checking by qualified technician. | Weekly check that cylinder at least ¼ full. | Spare cylinder available |
| Panic Buttons |  | Monthly check. |  |
| Printers |  |  | Spare cartridges |
| Scales |  | Compare with a known weight. |  |
| Smoke Alarms | Replace batteries annually. | Monthly check. | Spare battery |
| Spirometer | Servicing and calibration according to manufacturer’s recommendations. | Regular calibration using 3L syringe. Monthly clean or more frequent as required. | Disposable mouth pieces, preferably one way and preferably filtered |
| Steriliser | Annual servicing, calibration and validation by a qualified technician. | Maintenance and regular cleaning according to manufacturer’s instructions. | Deionised or Distilled water, bags or pouches, printer paper and ink cartridge, chemical indicators, sealing tape, spares for heat sealer if used |
| Sphygmomanometer | Annual checking against a mercury device or recently checked aneroid device. Cleaning and checking of mercury by qualified technician if required. |  | Cuffs various sizes Spare standard cuff and bladder, bulb and control valve |
| Telephone System |  | Quarterly check of battery back-up function. |  |

| Item | Servicing/Calibration | Monitoring/Checking | Parts/Supplies |
| --- | --- | --- | --- |
| Thermometer | Compare against mercury oral thermometer or other electronic thermometer. |  | Batteries, covers |
| UPS (Computer) |  | Quarterly monthly check of battery back-up function. |  |
| Vaccine Storage | Annual logging of fridge temperature using a calibrated data logger. Replace batteries on min/max thermometer annually. | Twice daily min/max temperature recording. | Spare batteries |